## Slide 1
KJM-MENA 4010 Module 2
Electrical measurements
With emphasis on fundamentals, electrical properties of materials and interfaces, and impedance spectroscopy

## Slide 2: Content/resources
Lectures
Demonstrations and excercises
Report
Compendium
Lecture handouts
Electronic resources

## Slide 3: Contents/goals – what we aim to learn
Electrical measurements – DC fundamentals
  Voltage, field, charge, current, power, energy
  Principles of voltage, current, and resistance measurements. Ohm’s law.
  Kirchhoff’s laws
  Internal resistance and load resistance
  2, 3, and 4 electrode measurements of resistance and conductance
Impedance spectroscopy
  Circuits of actual materials functions and discrete elements.
  Resistive/conductive, capacitive, inductive elements. Randles circuit. CPEs.
  Measurement, generation, and deconvolution of impedance spectra
  Impedance spectrometer: Main features and operation. Test discrete components.
High temperature electrical characterisation
  Sample preparation, geometries, specific properties. Shields, guards.
  Example setup – temperature, atmosphere, materials, contacts…
  Make and store measurements, example software
  Understand and deconvolute impedance spectra
  Understand and deconvolute temperature dependency of an example material
Thin film measurements

## Slide 4: Fundamentals Charge Current Potential Voltage

## Slide 5: Electrical charge
Charge is a physical property of matter that causes repulsive or attractive forces between objects of charge of the same or opposite sign.
Charge is quantized in multiples of e
The unit of electrical charge is C, coulomb
e = 1.602×10−19 C
As symbol of charge we use, for instance, q.
A proton has charge e
An electron has charge –e
Neutrons have charge 0
(Quarks have charges that are multiples of e/3)

## Slide 6: Current
Current I, is the flow of charge, is measured in ampere, A, and is measured in charge per second:
I = q/t
A = C/s
Current can be flow of for instance electronic charges in a metal wire or ionic charges in an electrolyte, or a mixture.
Current density i is current I per unit area a:
i = I / a
The unit of current density is A/m2 or A/cm2
DC Direct current – constant current (and voltage)
AC Alternating current – sinusoidal current (and voltage)

## Slide 7: Measurement of current and charge
Historically, current is measured by the electromagnetic force exerted on a magnet when the current flows in a coil. The force makes an indicator move on a current scale.
A modern ammeter measures the current electronically.
Ideally, the ammeter circuit has zero resistance so as not to affect the current. There is by necessity a small resistance.
Smaller current range, higher resistance. Why?
The ammeter is protected by a fuse. Check the fuse?
Don’t connect an ammeter directly to an electrical source: The fuse may blow if there is nothing that limits the current.
An ammeter may measure charge by integrating the current over time.

## Slide 8: Electric potential
The electric potential is defined and used in different ways
It is in one definition equal to the electric potential energy, namely the energy needed to add more charge
It thus has the unit joule per coulomb, J/C
The unit of electrical potential is volt V = J/C
As symbol for electric potential we use for instance φ (phi)

## Slide 9: Electric field and electrical field strength
An electric field is the region of space surrounding electrically charged objects
The electric field depicts the force exerted on other electrically charged objects
The field strength E is depicted by the density of field lines
E is the derivative of potential in space: E = -∇φ
E is the force exerted on a charged particle per unit charge: E = F/q
E has units of V/m or N/C

## Slide 10: Voltage
A voltage U is a difference in electric potential
U = Δφ = φ2 – φ1
The unit of voltage is the same as for potential: volt V
The potential energy Ep to take a unit charge from one potential to another is given by Ep = U*q
Hence U = Ep/q
The energy to rise an elementary charge e by 1 V is 1 electron volt, 1 eV
Voltage is electrical field strength multiplied with distance: U = -E*d

## Slide 11: Measure voltage
Historically, charge was first indicated and measured with gold leaf electroscopes. Voltage was later on measured with charge repulsive or attractive mechanical devices; electrometers.
Voltage can be measured by electromagnetic devices, similar to current, by driving current through the device.
Modern voltmeters use a range of electronic circuitry with high impedance field-effect transistors
Ideally, a voltmeter has high resistance and draws no current
  An electromagnetic indicator may have as little as 1 kΩ
  A cheap handheld analog multimeter may have 10 kΩ
  A handheld digital multimeter may have 1 MΩ
  A digital research multimeter may have 1 GΩ – 1 TΩ
  A good research multimeter may have > 1 TΩ (”electrometer”)

## Slide 12: Characteristics of various types of voltmeters
Electrometers
Coil-driven electromagnetic indicators
Electronic devices with transistors
Modern devices with MOSFET transistors

## Slide 13: Measure very high currents and voltages
Measure very high currents using a normal ammeter and a high current probe
  Clamps around wire
  Induction
  E.g. 1000 A
Measure very high voltages by using a normal voltmeter and a high voltage probe
  Voltage divider
  Safety
  E.g. 30 kV

## Slide 14: Materials and component properties: Resistivity and resistance Conductivity and conductance

## Slide 15: Resistivity and resistance
Charged particles in an electric field E feel a force F
The force sets up a net flux density and current density i
The ratio ρ (rho) = E/i is termed resistivity and is an intensive materials property
Resistivity has units (V/m)/(A/m2) = (V/A)m = ohm*m = Ωm
For an object we may instead express a current I and voltage U
The ratio R = U/I (Ohm’s law) is termed resistance and is an extensive property for the object
Resistance has units V/A = ohm = Ω
The resistance of a current-carrying object is obtained from the resistivity ρ, length l, and cross-sectional area a: R = ρ*l/a

## Slide 16: Conductivity and conductance
Conductivity σ (sigma) is the inverse of resistivity: σ = 1/ρ
Conductance G is the inverse of resistance: G = 1/R
The units for G and σ are S (siemens) and S/m, respectively.
(Other/older units for conductance comprise Ω-1, ohm-1, and mho)
G = σ*a/l

## Slide 17: Exercises
A rectangular solid sample has a length of 2 cm and a cross-section with sides 5 x 5 mm2. Electrodes for measurements are painted on its far faces.
  If its conductivity is 1000 S/cm, what is its conductance?
  And its resistance?
A circular disk has thickness 2 mm and diameter 2 cm. We paint electrodes on its two faces and measure the resistance.
  If the resistance is 10 Ω, what is the resistivity?
  If the conductance is 10 S, what is the conductivity?

## Slide 18: Total conductivity, transport numbers
The conductivity of a substance has contributions from all species, mechanisms, and pathways of charge carriers:
  Electronic and ionic
  Electronic: electrons and holes
  Ionic: cations and anions
  Or more detailed, for instance, protons, oxide ions, and metal cations
  Mechanisms: vacancies and interstitials
  Microstructural pathways: bulk, grain boundaries, surfaces…
The total conductivity is a sum of partial conductivities over all species, mechanisms, and pathways:
The fraction of the total conductivity (and ideally the fraction of any current going through the substance) is termed the transport number or transference number for s:

## Slide 19: Exercise
Normally, only one or two charge carriers, defects, mechanisms, or pathways dominate to the extent that we need to take them into account. The others can be neglected.
What dominates the conductance in Si? As-doped Si? Pt? NaCl(s)? NaCl(aq)? H2O(l)? HCl(aq)? Y-doped ZrO2? La2NiO4+δ? Alumina single crystal? Dense alumina ceramic? Porous alumina ceramic?

## Slide 20: Exercise
One can enhance or depress selected contributions for measurements or use
Discuss how you might affect the contributions below in the case of solid samples:
  Electronic conductivity
  Oxide ion conductivity
  Proton conductivity
  Bulk conductivity
  Grain boundary conductivity
  Outer surface conductivity
  Inner surface (pore wall) conductivity

## Slide 21: Series resistance contributions
Till now, we have looked at parallel possibilities that add to conductance and give more current
There are also many sources to series problems that add to resistance and give less current (more voltage):
  Bulk resistance
  Traps
  Grain boundary resistance
  Electrode (contact) resistance
Note the difference between grain boundary conduction and grain boundary resistance
What is the source of each one?
How can they be affected?

## Slide 22: Conductivity; charge, concentration, and mobility
The conductivity of a species s is given by its charge zs, volume concentration cs, and charge mobility us.
The charge is an integer multiple zs of e or F, depending of whether the concentration is given in number of particles or moles of particles per unit volume:
The concentration cs may arise from different models comprising doping and thermodynamics for electrons and/or point defects.
Charge mobility us is the product of mechanical mobility Bs and charge zse:

## Slide 23: Charge mobility; itinerant carriers (metallic mobility)
  In materials with metallic mobility (itinerant electrons or holes, broad bands) the mobility is determined by scattering, and the mobility is proportional to the mean free length between scattering events and inversely proportional to the electron or hole effective mass and the mean velocity at the mobile electrons’ energy level (Fermi level):
  Scatterers are defects (e.g. impurities) or phonons (lattice vibrations)
Both contribute to resistance in series:
Typical temperature dependencies:
Typically, impurities dominate at low T and lattice vibrations at high T.

## Slide 24: Charge mobility; diffusing carriers
  For ions that move by defects in materials and for non-itinerant (trapped) electrons in semiconductors, the mobility is determined by diffusion; thermally activated jumps from site to site:
  Note that usT (and thus σsT) is an exponential function of 1/T, and therefore the activation enthalpy may be extracted from the slope of a plot of ln(usT) or log(usT) versus 1/T (similar to an Arrhenius plot).
  Such electronic charge carriers are called small polarons – the electron deeply trapped in the relaxation of the lattice around itself. Small polaron mobilities are orders of magnitude smaller than itinerant (metallic) mobilities.
  Electronic charge carriers trapped in more shallow relaxations are called large polarons and have intermediate mobilities.

## Slide 25: …before we continue…Kirchhoff’s laws
Kirchhoff’s 1st law:
  The sum of currents flowing into a branching point is equal to the sum of currents flowing out in the branches.
  In other words, the total current equals the sum of currents in all parallel pathways; current is summed in parallel.
  Since the voltage over each parallel branch is the same, the current flowing in each is proportional to the conductance (or inversely proportional to the resistance) of that branch.
Kirchhoffs 2nd law:
  The sum of all potential changes (voltages) in a chosen direction around a closed circuit is zero.
  In other words, voltages are summed in series.
  Since the current is the same in all serial parts of the circuit, the voltage over each part is proportional to the resistance of that part.

## Slide 26: Resistance measurements
Historically, resistance was measured by Wheatstone bridges

## Slide 27: Modern voltage, current, and resistance measurements
Measure voltage
  Voltmeter (high input resistance) or Electrometer (very high input resistance)
Measure current
  In modern devices, current is measured by measuring the voltage over a standard resistor RS
Ohm’s law
  I = US/RS
Measure resistance
  Compare voltage UD over sample resistor RD with voltage US over standard resistor RS

## Slide 28: Internal resistance and load resistance
Voltage sources have effective internal resistance Ri
  From transport (ohmic) or reaction kinetics (electrochemical)
Voltage measurement devices have a finite inherent load resistance Rl .
  Big, but not infinitely big.
Delivery and measurement of voltage is only efficient and accurate if Rl >> Ri.
Modern MOSFET transistors ensure high Rl in all devices.

## Slide 29: 2, 3, and 4 electrode measurements of resistance and conductance
Current flows through the voltage source of the measuring device, the standard resistor Rs, the sample and its electrodes, and all connections in between.
The device measures voltage using the same 2 terminals or through separate terminals, so as to have 4 in total.
The voltage circuit itself is of high impedance, has virtually no current, and introduces no voltage drop.
The voltage drop and hence the calculated resistance/conductance is between where the voltage contacts meet the current circuit.
By placing them nearer the sample, you can eliminate voltage drops (resistance) in connections, wires, and one or both electrodes.

## Slide 30: Impedance spectroscopy

## Slide 31: Impedance spectroscopy (IS) (= Electrochemical impedance spectroscopy (EIS))
  Name stems from historical use of oscilloscopes to record amplitude and phase
  Better name today might be impedance spectrometry
  Use AC current and voltage.
  Obtain impedance (ratio) and phase angle between voltage and current.
  Vary the AC frequency over many orders of magnitude and fit amplitude and phase angle to an assumed model equivalent circuit mimicking the physical sample
  Obtain individual component parameters values where possible

## Slide 32: AC voltage and current. Impedance and power.
AC voltage
the amplitude can be expressed as
  Amplitude U0
  Peak-to-peak voltage, Up-p = 2U0
  Root mean square (rms) voltage
  Urms = U0/2 = Up-p/(22)
Resulting AC current can have a different amplitude and phase angle as a result of the impedance:
Impedance Z* = U* / I*
Instantaneous power P = U* I*
Power integrated over time:

## Slide 33: Complex impedance and admittance
Z* = U*/I* = Z/ + jZ// = R + jX
  Z* Complex impedance
  |Z| Impedance
  R Resistance
  X Reactance
Y* = 1/Z* = I*/U* = Y/ + jY// = G + jB
  Y* Complex impedance
  |Y| Admittance
  G Conductance
  B Susceptance

## Slide 34: Impedance and admittance of discrete elements
Series RCL
Z* = Z/ + jZ// = R + jX
Parallel (RCL)
Y* = Y/ + jY// = G + jB
Resonance for (RC) when G = ωC or C=G/ω or ω=1/(RC)

## Slide 35: Y*=1/Z* and Z*=1/Y*

## Slide 36: Circuits of actual materials functions and of discrete elements.
Introductory warm-up:
  The circuit contains discrete elements representing physical properties of a typical ceramic mixed ionic electronic conductor
Upper rail: Electronic transport rail
  Re Electronic conduction (NOTE: No electrode charge transfer resistance assumed for electrons)
Lower rail: Ionic transport rail
  Rb ionic bulk conduction (grain interior)
  Cg geometric capacitance of the material
  Rgb grain boundary resistance
  Qgb grain boundary capacitance element
  Rct Electrode red-ox reaction charge transfer
  Cdl Double layer capacitance
  Rd Diffusion resistance
  Qd Diffusion-related (chemical) storage
L: inductive element
  From sample, wires, and instrument parasitics
Draw a sample with microstructure and place the elements
What are the assumptions?
How does everything add up?
Microstructure
Equivalent Circuit

## Slide 37: Mathematics of some of the parameters
Bulk conductivity
  Linear property
  Sum of partial conductivity from each charge carrier
  Partial conductivity: Charge, concentration, and mobility
Grain boundary transport
  Linear property
  Dominated by space charge layers
Electrode impedance
  Butler-Volmer theory – linear case
  Tafel behaviour – non-linear case

## Slide 38: Conduction takes many paths
Bulk (grain interior)
Grain boundaries
  Series resistance
  Parallel conductance
  Core region
  Space charge layers
Enhanced transport on external and internal surfaces
Transport in adsorbed water
Electrode impedance

## Slide 39: Bulk and grain boundary transport..
in series (transport across) can be delineated by impedance spectroscopy;
in parallel (transport along) can not be delineated by impedance spectroscopy.

## Slide 40: Resistive/conductive, capacitive, inductive elements
Resistor (conductor) R=1/G; no phase shift between I and U
Capacitor
Inductor L

## Slide 41: Constant phase elements (CPE)
Constant phase elements (CPE); symbol Q
Mathematically mimics the impedance of real, non-ideal, processes and components; dispersions in time constants
Q has admittance
n = 1 Capacitor
n = 0 Conductor/Resistor
n = -1 Inductance
n = 0.5 Warburg element
Values in between: Intermediate behaviours

## Slide 42: Electrodes

## Slide 43: Aqueous and solid state electrolyte electrodes

## Slide 44: Randles circuits
Charge transfer (ct)
Double layer (dl)
Diffusion (d) and chemical storage
Non-blocking Semi-blocking or Blocking

## Slide 45: Electrode charge transfer kinetics: Butler-Volmer (BV)
An electrode has an equilibrium Nernst potential
  (vs another electrode)
At this potential there is no net current; i = 0.
  Forward and backward currents i0 are equal and cancel
Other potentials can be represented as an overpotential η
Current vs overpotential according to the BV equation:

## Slide 46: At small overpotentials, we can linearise BV

## Slide 47: At large overpotentials, we can neglect the backward direction

## Slide 48: Mass transport limitations
η-i curve with mass transport limiting current il (left)
Typical fuel cell discharge i-U curve with activation or concentration loss, ohmic loss, and mass transport limiting current (right)

## Slide 49: Electrode impedances; DC or AC techniques
DC
  Chronovoltammetry: Steady state: One potential at the time
  Linear sweep voltammetry (LSV): Very slow scan
    Probing U/I of the resistances (kinetics)
Cyclic voltammetry
  Faster, usually cyclic, voltammetry
    Probing U/I at one time constant; info about R’s and C’s
AC impedance spectroscopy
  Probing all time constants at one U or I; deconvolute circuit
Other: Noise+FT. Current interruption. Thermal depolarisation...

## Slide 50: Measurement, generation, and deconvolution of impedance spectra
Impedance spectra are measured logarithmically
  Often from high to low frequency.
We can generate impedance spectra from equivalent circuits
We can deconvolute impedance spectra
  Generate partial circuits
    fit and subtracted one by one while building up the circuit
    Final non-linear least squares (NLLS) fitting of the whole circuit. Use software “Equivalent Circuit” (EQ) or similar.
  Notation: Each level of parenthesis goes from series to parallel. The start is series by default.

## Slide 51: Geometry and parasitics

## Slide 52: Geometry
Bulk: Correct for area and thickness to obtain materials specific conductivity from conductance
Electrodes: Correct for area to obtain area specific conductivity or resistivity
Grain boundaries: Use the brick layer model (BLM)

## Slide 53: Parasitics
Parallel and series
Typical series elements
  Rparasitic 0.1 ohm
  Lparasitic 10-6 H
Typical parallel elements
  Gparasitic 10-6 S (106 ohm)
  Cparasitic 10-12 F

## Slide 54: Shielding and guarding
Shields and guards are connected to ground (earth) levels of the electronics and prevent fields and currents from passing through them.
Used to reduce noise and parasitics
  Three shield systems Guard principle Surface guard
  Note: Some impedance spectrometers need the shields to be connected together, as they are used for countercurrent compensation

## Slide 55: Electrometer and driven shields
High impedance samples and cells give noise and parasitic loads.
Shields can shield/guard against unwanted noise
But they add parasitic load to high impedance samples
The solutions:
High input impedance preamplifier.
  Long wires operated at high signal to noise ratio and low impedance
Driven shields
  Shield placed at same potential as line – no load current can flow from line to shield.
  These are parts of modern electrometers and potentiostats

## Slide 56: Measurements of thin films
Conducting substrate
  Measure across the film
  Film should be a poor conductor or dielectric
  Parasitics
    Effect of space charge
    Series resistance
      Wires
      Spreading resistance
Insulating substrate
  Measure along the film
  Film should be a good conductor
  Parasitics
    Surface conduction
    Substrate conduction and capacitance
  Van der Pauw geometry may be used

## Slide 57: Impedance spectrometer – manual operation excercise 1
Make a circuit (RC)
Calculate its resonance frequency
Connect it to the impedance spectrometer
Learn to control
  frequency, potentiostatic (U) or galvanostatic (I) control, Uosc, Ubias, measurement speed and accuracy, series vs parallel measurement and representation
Measure (display) at three different frequencies:
  Impedance and phase angle
  Admittance and phase angle
  R and X, R and C in parallel and series mode
  G and B, G and C in parallel and series mode
Do the variations make sense?

## Slide 58: Impedance spectrometer – manual operation excercise 2
Connect a triangle of equally big resistors (e.g. 1 kohm each)
Write its equivalent circuit, assuming two of the connection points as terminals
Calculate the resistance between the two terminal points.
Measure it and compare.
Now, additionally connect the third terminal to shield or ground.
This eliminates the pathway passing the third terminal.
Measure the resistance again between the same two terminals, with the third guarded.
Does this fit with the measurement?

## Slide 59: High temperature ramp
Sample: 8 mol% Y2O3-doped ZrO2 (8YSZ), pressed and sintered. Pt paste + mesh electrodes applied and fired.
Measure the outer dimensions of the sample and of the electrodes.
Set up in a ProboStat cell using alumina and Pt parts. Supply wet air.
Measurements using an impedance spectrometer at 1 V rms: 4 wires, 2 electrodes. Be sure to set the ProboStat switches correctly.
(If the sample has a third electrode as a ring, face it down onto a ring electrode contact. Let this be available at one of the electrode outlets, preferably LV.)
Measure from e.g. 1100°C to room temperature.
Collect conductivity at 10 Hz, 1000 Hz, and 100 kHz.
Plot log(σT) vs 1/T for one of them. Report the main activation energies.
For concave curves, try to fit the total curve to a sum of two straight lines.
Repeat ramp using dry instead of wet air.
At room temperature, measure with and without guard ring electrode connected to ground (shields). Make a ramp with the guard ring connected.

## Slide 60: High temperature impedance spectroscopy
Repeat the above (without guard), but for each 100°C equilibrate and take an impedance spectrum (e.g. 1 MHz – 0.01 Hz).
Deconvolute according to a simple circuit that can explain the features of the spectrum.
Write a short report showing the temperature ramp and the impedance spectrum at one temperature, and provide a brief explanation and the parameters obtained.
Include a brief discussion of the most important sources of error and uncertainty involved.
(Not part of the report.) At the end, the group may gather data from each deconvolution in order to create a plot of individual parameters vs 1/T. Compare it with the ramps at fixed frequency.